## Cover
Things to look into
Electron energy baselines
MPC enhanced couch full range of motion 0.78 mm
2.5 MV beam
CBCT HU calibration

## Dosimetry
Dosimetry
1.1 X-ray profile change from baseline (TG142: 1%)
Beam Data acquired during commissioning is used as baseline. Located at:
With ICE Profiler, capture profiles using the notated buildup, wedge plates, etc.
NOTE: Depth in file name includes 0.9cm inherent buildup. So solid water depth is: File Name Depth - 0.9cm = Solid Water cm
NOTE: Some Beam data acquired at different depths than annual measurement. This is for simplification, and % Differences will serve as Baseline Differences
Buildup Profiler files (photons):
100 cm SSD
Field Size: Variable
| Energy | Field Size | Depth (cm) | 2020 Baseline |  |  |  | ICE Profiler Measured |  |  |  | % Difference
|  |  |  | Y Axis (Inplane) |  | X Axis (Crossplane) |  | Y Axis (Inplane) |  | X Axis (Crossplane) |  | Y Axis (Inplane) |  | X Axis (Crossplane)
|  |  |  | Flatness (%) | Symmetry (%) | Flatness (%) | Symmetry (%) | Flatness (%) | Symmetry (%) | Flatness (%) | Symmetry (%) | Flatness (%) | Symmetry (%) | Flatness (%) | Symmetry (%)
| 6X | 4x4 | 1.4 | 6.8 | 0.2 | 12.8 | 0.2 | [10 empty] | Profiler Analysis Setup >> Setup > Analysis
|  | 10x10 |  | 1.1 | 0.2 | 1.3 | 0.2
|  | 20x20 |  | 1.5 | 0.4 | 1.3 | 0.2
|  | 4x4 | 4.9 | 7.7 | 0.6 | 12.8 | 0.4
|  | 10x10 |  | 2 | 0.2 | 2.2 | 0.3
|  | 20x20 |  | 1.4 | 0.3 | 1.1 | 0.2
|  | 4x4 | 9.9 | 8.8 | 1.3 | 10.1 | 0.1
|  | 10x10 |  | 2.6 | 0.2 | 2.9 | 0.2
|  | 20x20 |  | 2.6 | 0.3 | 2.3 | 0.1 | 3.1 | 0.9 | 2.5 | -0.4 | -0.5 | -0.6 | -0.2 | 0.5
| 6FFF | 4x4 | 1.4 | 7.3 | 0 | 13.9 | -0.5
|  | 10x10 |  | 4.7 | 0.2 | 5.1 | -0.1
|  | 20x20 |  | 11 | 0.4 | 11.2 | -0.3
|  | 4x4 | 4.9 | 8.1 | 0.4 | 13.3 | 0.2
|  | 10x10 |  | 5.9 | 0.4 | 6 | 0
|  | 20x20 |  | 11.8 | 0.5 | 11.9 | -0.3
|  | 4x4 | 9.9 | 9.1 | 1.1 | 10.6 | 0
|  | 10x10 |  | 6.5 | 0.3 | 6.6 | -0.1
|  | 20x20 |  | 12.8 | 0.6 | 12.8 | -0.3 | 12.6 | 0.6 | 12.1 | 0 | 0.2 | 0 | 0.7 | -0.3
| 10X | 4x4 | 1.4 | 7.1 | 0.3 | 13.5 | 0.2
|  | 10x10 |  | 1 | 0.3 | 1.4 | 0.6
|  | 20x20 |  | 1.5 | 0.2 | 1.7 | 0.6
|  | 4x4 | 4.9 | 8.4 | 0.5 | 13.6 | 0.6
|  | 10x10 |  | 2.1 | 0.3 | 2.2 | 0.6
|  | 20x20 |  | 1.1 | 0.3 | 1.2 | 0.5
|  | 4x4 | 9.9 | 9.5 | 1.2 | 11.1 | 0.4
|  | 10x10 |  | 2.4 | 0.3 | 2.6 | 0.5
|  | 20x20 |  | 2 | 0.3 | 1.9 | 0.4 | 2.3 | 0.9 | 1.6 | -0.1 | -0.3 | -0.6 | 0.3 | 0.5
| 10FFF | 4x4 | 1.4 | 9.2 | 0.5 | 14.7 | -0.6
|  | 10x10 |  | 9 | 0.6 | 9.3 | 0
|  | 20x20 |  | 19.4 | 0.9 | 19.5 | -0.2
|  | 4x4 | 4.9 | 10 | 0.7 | 14.4 | 0.1
|  | 10x10 |  | 9.9 | 0.8 | 9.9 | 0
|  | 20x20 |  | 20 | 1 | 20.1 | -0.3
|  | 4x4 | 9.9 | 11 | 1.5 | 12 | 0.1
|  | 10x10 |  | 10.4 | 1 | 10.3 | -0.1
|  | 20x20 |  | 20.5 | 1.2 | 20.5 | -0.2 | 19.3 | 0.8 | 19 | 0.3 | 1.2 | 0.4 | 1.5 | -0.5
| 18X | 4x4 | 3.2 | 9.3 | 0.8 | 14.7 | 0.1
|  | 10x10 |  | 2.1 | 0.3 | 2.1 | 0.3
|  | 20x20 |  | 1 | 0.1 | 1.1 | 0.3
|  | 4x4 | 4.9 | 10 | 1 | 14.5 | 0.3
|  | 10x10 |  | 2.4 | 0.3 | 2.3 | 0.3
|  | 20x20 |  | 1.1 | 0.1 | 1.2 | 0.2
|  | 4x4 | 9.9 | 10.7 | 1.6 | 12.2 | 0.3
|  | 10x10 |  | 2.5 | 0.3 | 2.4 | 0.3
|  | 20x20 |  | 1.3 | 0.1 | 1.1 | 0.2 | 1.9 | 1.4 | 1.1 | -0.2 | -0.6 | -1.3 | 0 | 0.4
1.2 Electron profile change from baseline (TG142: 1%)
Beam Data acquired during commissioning is used as baseline. Located at:
With ICE Profiler, capture profiles using the notated buildup, wedge plates, etc.
NOTE: Depth in file name includes 0.9cm inherent buildup. So solid water depth is: File Name Depth - 0.9cm = Solid Water cm
Wedge Plate Profiler files (electrons):
100 cm SSD
25x25 Cone | Plastic Wedge Plate
| Energy, Acquisition Parameters | Measured |  |  |  | Baseline |  |  |  | % Difference
|  | Y Axis (Inplane) |  | X Axis (Crossplane) |  | Y Axis (Inplane) |  | X Axis (Crossplane) |  | Y Axis (Inplane) |  | X Axis (Crossplane)
|  | Flatness (%) | Symmetry (%) | Flatness (%) | Symmetry (%) | Flatness (%) | Symmetry (%) | Flatness (%) | Symmetry (%) | Flatness (%) | Symmetry (%) | Flatness (%) | Symmetry (%)
| 6 MeV, Wedge Plate, 25x25 Cone |  |  |  |  | 1.3 | -0.1 | 1.5 | 0.2
| 9 MeV, Wedge Plate, 25x25 Cone |  |  |  |  | 1.7 | -0.4 | 1.5 | 0.4
| 12 MeV, Wedge Plate, 25x25 Cone |  |  |  |  | 1.4 | -0.1 | 1.3 | 0.3
| 15 MeV, Wedge Plate, 25x25 Cone |  |  |  |  | 1.6 | -0.3 | 1.5 | 0.4
| 18 MeV, Wedge Plate, 25x25 Cone |  |  |  |  | 2 | -0.6 | 1.7 | 0.3
1.3 SRS arc rotation mode (TG142: See Below)
Monitor Units Set vs Delivered: 1.0 MU or 2% (whichever is greater)
Gantry Arc Set vs Delivered: 1.0o or 2% (whichever is greater)
1.4 TG-51 X-ray/Electron output calibration (TG142: 1% absolute)
See "TG-51 Photon" and "TG-51 Electron" Tabs
1.5 X-ray beam quality (TG142: 1% from baseline)
Varian GBD, validated against measured data, provides baseline
Measured data taken from TG51 analysis
PDDs stored at:
|  | 6 MV |  |  | 6FFF |  |  | 10 MV |  |  | 10FFF |  |  | 18 MV
|  | GBD | Measured | % Diff | GBD | Measured | % Diff | GBD | Measured | % Diff | GBD | Measured | % Diff | Commis. | Measured | % Diff
| PDD10(%) | 66.4 | 67.1107 | 1.0703 | 63.5 | 64.0934 | 0.93456 | 73.5 | 73.9308 | 0.58606 | 71.1 | 71.0579 | 0.059235 | 79.5 | 79.8643 | 0.45818
1.6 Electron beam quality (TG142: 1mm)
Varian GBD, validated against measured data, provides baseline
Measured data taken from TG51 analysis
PDDs stored at:
|  | 6 MeV |  |  | 9 MeV |  |  | 12 MeV |  |  | 16 MeV |  |  | 20 MeV
|  | GBD | Measured | Diff (mm) | GBD | Measured | Diff (mm) | GBD | Measured | Diff (mm) | GBD | Measured | Diff (mm) | GBD | Measured | Diff (mm)
| R50 (cm) | 2.36 | 2.351 | 0.090167 | 3.6 | 3.6086 | 0.086175 | 4.98 | 5.0769 | 0.96899 | 6.61 | 6.7799 | 1.6991 | 8.25 | 8.4136 | 1.636
1.7 Spot check of field size dependent output factors for X-Ray (TG142: See Below)
Field Size <4 x 4 cm2: 2%; Field Size ≥4x4 cm2: 1%
Eclipse provides baseline data. Patient: Annual, QA
SSD = 100cm |  | 100 MU
Depth = 10cm |  | 600 MU/min (Unless otherwise indicated)
|  | 6 MV | [9 empty] | 6 FFF
| Field Size: | 5 | 8 | 10 | 12 | 15 | 20 | 25 | 30 |  | Field Size: | 5 | 8 | 10 | 12 | 15 | 20 | 25 | 30
| Readings (nC) | [9 empty] | Readings (nC)
R2
R3
| Average | [9 empty] | Average
| OF | [9 empty] | OF
| Eclipse | 0.895 | 0.964 | 1 | 1.027 | 1.06 | 1.101 | 1.132 | 1.153 |  | Eclipse | 0.90424 | 0.9686 | 1 | 1.0235 | 1.0502 | 1.0816 | 1.102 | 1.1162
| % Diff | [9 empty] | % Diff
|  | 10 MV | [9 empty] | 10 FFF
| Field Size: | 5 | 8 | 10 | 12 | 15 | 20 | 25 | 30 |  | Field Size: | 5 | 8 | 10 | 12 | 15 | 20 | 25 | 30
| Readings (nC) | [9 empty] | Readings (nC)
| Average | [9 empty] | Average
| OF | [9 empty] | OF
| Eclipse | 0.91304 | 0.97011 | 1 | 1.0204 | 1.0462 | 1.0774 | 1.1005 | 1.1168 |  | Eclipse | 0.93961 | 0.98174 | 1 | 1.014 | 1.0309 | 1.0478 | 1.059 | 1.0674
| % Diff | [9 empty] | % Diff
18 MV
| Field Size: | 5 | 8 | 10 | 12 | 15 | 20 | 25 | 30
Readings (nC)
Average
OF
| Eclipse | 0.9225 | 0.975 | 1 | 1.0162 | 1.0337 | 1.0588 | 1.0775 | 1.0913
% Diff
1.8 Output factors for electron applicators (TG142: 2% from baseline)
Beam Config. provides baseline data.
100 cm SSD
| d = 1.3 cm | 6 MeV | [5 empty] | d = 2.0 cm | 9 MeV | [5 empty] | d = 2.8 cm | 12 MeV
| Field Size: | 6x6 | 10x10 | 15x15 | 20x20 | 25x25 |  | Field Size: | 6x6 | 10x10 | 15x15 | 20x20 | 25x25 |  | Field Size: | 6x6 | 10x10 | 15x15 | 20x20 | 25x25
| Readings (nC) | [6 empty] | Readings (nC) | [6 empty] | Readings (nC)
| Average | [6 empty] | Average | [6 empty] | Average
| OF | [6 empty] | OF | [6 empty] | OF
| Eclipse | 0.949 | 1 | 1.001 | 1.044 | 1.067 |  | Eclipse | 0.99 | 1 | 0.995 | 1.01 | 1.011 |  | Eclipse | 0.961 | 1 | 1.004 | 1.004 | 0.99
| % Diff | [6 empty] | % Diff | [6 empty] | % Diff
| d = 2.8 cm | 16 MeV | [5 empty] | d = 2.0 cm | 20 MeV
| Field Size: | 6x6 | 10x10 | 15x15 | 20x20 | 25x25 |  | Field Size: | 6x6 | 10x10 | 15x15 | 20x20 | 25x25
| Readings (nC) | [6 empty] | Readings (nC)
| Average | [6 empty] | Average
| OF | [6 empty] | OF
| Eclipse | 0.977 | 0.999 | 0.993 | 0.989 | 0.977 |  | Eclipse | 1.001 | 0.999 | 0.978 | 0.972 | 0.957
| % Diff | [6 empty] | % Diff
1.9 X-ray monitor unit linearity (TG142: 2-4 MU: 5%; ≥5 MU: 2%)
10cm depth | Farmer chamber 9244
600 MU/min | 100 cm SSD 30cm water tank
|  | 6 MV |  |  | 6 FFF |  |  | 10 MV |  |  | 10 FFF |  |  | 18 MV
| MU | Reading (nC) | nC/MU | % Diff from 100 MU | Reading (nC) | nC/MU | % Diff from 100 MU | Reading (nC) | nC/MU | % Diff from 100 MU | Reading (nC) | nC/MU | % Diff from 100 MU | Reading (nC) | nC/MU | % Diff from 100 MU
| 2 | 0.28 | 0.14 | 0.21475 | 0.2635 | 0.13175 | 2.609 | 0.305 | 0.1525 | 0.85979 | 0.277 | 0.1385 | 0.072254 | 0.328 | 0.164 | 1.3597
| 5 | 0.699 | 0.1398 | 0.071582 | 0.64 | 0.128 | 0.31153 | 0.756 | 0.1512 | 0 | 0.695 | 0.139 | 0.43353 | 0.814 | 0.1628 | 0.61805
| 50 | 6.958 | 0.13916 | 0.38654 | 6.413 | 0.12826 | 0.10903 | 7.555 | 0.1511 | 0.066138 | 6.919 | 0.13838 | 0.014451 | 8.096 | 0.16192 | 0.074166
| 100 | 13.97 | 0.1397 | 0 | 12.84 | 0.1284 | 0 | 15.12 | 0.1512 | 0 | 13.84 | 0.1384 | 0 | 16.18 | 0.1618 | 0
| 500 | 69.61 | 0.13922 | 0.34359 | 64.13 | 0.12826 | 0.10903 | 75.52 | 0.15104 | 0.10582 | 69.1 | 0.1382 | 0.14451 | 80.97 | 0.16194 | 0.086527
1.10 Electron monitor unit linearity (TG142: ≥5 MU, 2%)
600 MU/min | 10x10 cone
|  | 6 MeV @ 1.3cm |  |  | 9 MeV @ 2cm |  |  | 12 MeV @ 2.8 cm |  |  | 16 MeV @ 2.8cm |  |  | 20 MeV @ 2 cm
| MU | Reading (nC) | nC/MU | % Diff from 100 MU | Reading (nC) | nC/MU | % Diff from 100 MU | Reading (nC) | nC/MU | % Diff from 100 MU | Reading (nC) | nC/MU | % Diff from 100 MU | Reading (nC) | nC/MU | % Diff from 100 MU
| 5 | 0.983 | 0.1966 | 0.51125 | 0.9955 | 0.1991 | 0.91232 | 1 | 0.2 | 0.3009 | 1.0165 | 0.2033 | 0.79326 | 1.038 | 0.2076 | 0.7278
| 50 | 9.788 | 0.19576 | 0.0818 | 9.848 | 0.19696 | 0.17233 | 9.95 | 0.199 | 0.2006 | 10.09 | 0.2018 | 0.049579 | 10.3 | 0.206 | 0.04852
| 100 | 19.56 | 0.1956 | 0 | 19.73 | 0.1973 | 0 | 19.94 | 0.1994 | 0 | 20.17 | 0.2017 | 0 | 20.61 | 0.2061 | 0
| 500 | 97.76 | 0.19552 | 0.0409 | 98.36 | 0.19672 | 0.29397 | 99.43 | 0.19886 | 0.27081 | 100.8 | 0.2016 | 0.049579 | 103 | 0.206 | 0.04852
1.11 X-ray output constancy vs dose rate (TG142: 2%)
50 MU | Farmer chamber 9244
10 cm depth | 100 cm SSD 30cm water tank
|  | 6 MV |  | 6 FFF |  | 10 MV |  | 10 FFF |  | 18 MV
| MU/min | Reading (nC) | % Diff from 600 MU/min | Reading (nC) | % Diff from 1400 MU/min | Reading (nC) | % Diff from 600 MU/min | Reading (nC) | % Diff from 2400 MU/min | Reading (nC) | % Diff from 600 MU/min
| 100 | 6.976 | 0.2587 |  |  | 7.569 | 0.10581 |  |  | 8.099 | 0.19795
| 200 | 6.967 | 0.12935 |  |  | 7.556 | 0.066129 |  |  | 8.093 | 0.12372
| 300 | 6.964 | 0.086232 |  |  | 7.555 | 0.079355 |  |  | 8.087 | 0.049487
| 400 | 6.96 | 0.028744 | 6.423 | 0.031148 | 7.556 | 0.066129 | 6.932 | 0.28935 | 8.094 | 0.13609
| 500 | 6.963 | 0.07186 |  |  | 7.555 | 0.079355 |  |  | 8.091 | 0.098973
| 600 | 6.958 | 0 | 6.417 | 0.062296 | 7.561 | 0 |  |  | 8.083 | 0
| 800 |  |  | 6.424 | 0.046722 |  |  | 6.925 | 0.18808
| 1000 |  |  | 6.411 | 0.15574
| 1200 |  |  | 6.42 | 0.015574 |  |  | 6.91 | 0.028935
| 1400 |  |  | 6.421 | 0
| 1600 | [6 empty] | 6.92 | 0.11574
| 2000 | [6 empty] | 6.911 | 0.014468
| 2400 | [6 empty] | 6.912 | 0
1.12 Electron output constancy vs dose rate (TG142: 2%)
50 MU | 10x10 cone
|  | 6 MeV @ 1.3cm |  | 9 MeV @ 2cm |  | 12 MeV @ 2.8 cm |  | 16 MeV @ 2.8cm |  | 20 MeV @ 2cm
| MU/min | Reading (nC) | % Diff from 600 MU/min | Reading (nC) | % Diff from 600 MU/min | Reading (nC) | % Diff from 600 MU/min | Reading (nC) | % Diff from 600 MU/min | Reading (nC) | % Diff from 600 MU/min
| 100 | 9.801 | 0.071472 | 9.858 | 0.65346 | 9.966 | 0.15074 | 10.11 | 0.19822 | 10.32 | 0
| 200 | 9.793 | 0.01021 | 9.849 | 0.56157 | 9.958 | 0.070345 | 10.09 | 0 | 10.31 | 0.096899
300
| 400 | 9.786 | 0.081683 | 9.85 | 0.57178 | 9.958 | 0.070345 | 10.08 | 0.099108 | 10.3 | 0.1938
500
| 600 | 9.794 | 0 | 9.844 | 0.51052 | 9.951 | 0 | 10.09 | 0 | 10.32 | 0
700
| 800 | 9.783 | 0.11231 | 9.857 | 0.64325 | 9.963 | 0.12059 | 10.09 | 0 | 10.3 | 0.1938
900
| 1000 | 9.788 | 0.061262 | 9.848 | 0.55136 | 9.95 | 0.010049 | 10.09 | 0 | 10.3 | 0.1938
1.12 X-ray output constancy vs gantry angle (TG142: 1% from baseline)
1.13 Electron output constancy vs gantry angle (TG142: 1% from baseline)
1.14 Electron and x-ray off-axis factor constancy vs gantry angle (TG142: 1% from baseline)
Photons
| QA3 mounted in IMF, 20x20cm2, 100 MU at gantry angles G=270, 0, 90, 180 | [7 empty] | RATIO OF OUTER CHAMBER TO CENTRAL CHAMBER | [7 empty] | % CHANGE FROM 0 DEGREES | [8 empty] | SYMMETRY |  | % CHANGE FROM 0 DEGREES
| Gantry angle | Energy | Center | Top (1) | Right (2) | Bottom (3) | Left (4) |  | Gantry angle | Energy | Center | Top (1) | Right (2) | Bottom (3) | Left (4) |  | Gantry angle | Energy | Center | Top | Right | Bottom | Left | Gantry angle | Energy | Chamber 1/3 | Chamber 2/4 | Chamber 1/3 | Chamber 2/4
| 0 degrees | 6 MV | [6 empty] | 0 degrees | 6 MV | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! |  | 0 degrees | 6 MV | 0 | 0 | 0 | 0 | 0 | 0 degrees | 6 MV | #DIV/0! | #DIV/0! | 0 | 0
|  | 6FFF | [7 empty] | 6FFF | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! |  |  | 6FFF | 0 | 0 | 0 | 0 | 0 |  | 6FFF | #DIV/0! | #DIV/0! | 0 | 0
|  | 10MV | [7 empty] | 10MV | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! |  |  | 10MV | 0 | 0 | 0 | 0 | 0 |  | 10MV | #DIV/0! | #DIV/0! | 0 | 0
|  | 10FFF | [7 empty] | 10FFF | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! |  |  | 10FFF | 0 | 0 | 0 | 0 | 0 |  | 10FFF | #DIV/0! | #DIV/0! | 0 | 0
|  | 18 MV | [7 empty] | 18 MV | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! |  |  | 18 MV | 0 | 0 | 0 | 0 | 0 |  | 18 MV | #DIV/0! | #DIV/0! | 0 | 0
| 90 degrees | 6 MV | [6 empty] | 90 degrees | 6 MV | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! |  | 90 degrees | 6 MV | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | 90 degrees | 6 MV | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
|  | 6FFF | [7 empty] | 6FFF | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! |  |  | 6FFF | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! |  | 6FFF | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
|  | 10MV | [7 empty] | 10MV | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! |  |  | 10MV | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! |  | 10MV | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
|  | 10FFF | [7 empty] | 10FFF | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! |  |  | 10FFF | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! |  | 10FFF | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
|  | 18 MV | [7 empty] | 18 MV | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! |  |  | 18 MV | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! |  | 18 MV | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
| 180 degrees | 6 MV | [6 empty] | 180 degrees | 6 MV | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! |  | 180 degrees | 6 MV | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | 180 degrees | 6 MV | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
|  | 6FFF | [7 empty] | 6FFF | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! |  |  | 6FFF | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! |  | 6FFF | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
|  | 10MV | [7 empty] | 10MV | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! |  |  | 10MV | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! |  | 10MV | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
|  | 10FFF | [7 empty] | 10FFF | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! |  |  | 10FFF | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! |  | 10FFF | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
|  | 18 MV | [7 empty] | 18 MV | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! |  |  | 18 MV | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! |  | 18 MV | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
| 270 degrees | 6 MV | [6 empty] | 270 degrees | 6 MV | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! |  | 270 degrees | 6 MV | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | 270 degrees | 6 MV | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
|  | 6FFF | [7 empty] | 6FFF | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! |  |  | 6FFF | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! |  | 6FFF | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
|  | 10MV | [7 empty] | 10MV | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! |  |  | 10MV | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! |  | 10MV | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
|  | 10FFF | [7 empty] | 10FFF | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! |  |  | 10FFF | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! |  | 10FFF | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
|  | 18 MV | [7 empty] | 18 MV | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! |  |  | 18 MV | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! |  | 18 MV | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
Electrons
| QA3 mounted in IMF, 20x20cm2 (No Cone), 100 MU, at gantry angles G=270, 0, 90, 180 | [7 empty] | RATIO OF OUTER CHAMBER TO CENTRAL CHAMBER | [7 empty] | % CHANGE FROM 0 DEGREES | [8 empty] | SYMMETRY |  | % CHANGE FROM 0 DEGREES
| Gantry angle | Energy | Center | Top (1) | Right (2) | Bottom (3) | Left (4) |  | Gantry angle | Energy | Center | Top (1) | Right (2) | Bottom (3) | Left (4) |  | Gantry angle | Energy | Center | Top (1) | Right (2) | Bottom (3) | Left (4) | Gantry angle | Energy | Chamber 1/3 | Chamber 2/4 | Chamber 1/3 | Chamber 2/4
| 0 degrees | 6 MeV | [6 empty] | 0 degrees | 6 MeV | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! |  | 0 degrees | 6 MeV | 0 | 0 | 0 | 0 | 0 | 0 degrees | 6 MeV | #DIV/0! | #DIV/0! | 0 | 0
|  | 9 MeV | [7 empty] | 9 MeV | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! |  |  | 9 MeV | 0 | 0 | 0 | 0 | 0 |  | 9 MeV | #DIV/0! | #DIV/0! | 0 | 0
|  | 12 MeV | [7 empty] | 12 MeV | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! |  |  | 12 MeV | 0 | 0 | 0 | 0 | 0 |  | 12 MeV | #DIV/0! | #DIV/0! | 0 | 0
|  | 16 MeV | [7 empty] | 16 MeV | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! |  |  | 16 MeV | 0 | 0 | 0 | 0 | 0 |  | 16 MeV | #DIV/0! | #DIV/0! | 0 | 0
|  | 20 MeV | [7 empty] | 20 MeV | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! |  |  | 20 MeV | 0 | 0 | 0 | 0 | 0 |  | 20 MeV | #DIV/0! | #DIV/0! | 0 | 0
| 90 degrees | 6 MeV | [6 empty] | 90 degrees | 6 MeV | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! |  | 90 degrees | 6 MeV | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | 90 degrees | 6 MeV | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
|  | 9 MeV | [7 empty] | 9 MeV | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! |  |  | 9 MeV | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! |  | 9 MeV | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
|  | 12 MeV | [7 empty] | 12 MeV | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! |  |  | 12 MeV | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! |  | 12 MeV | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
|  | 16 MeV | [7 empty] | 16 MeV | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! |  |  | 16 MeV | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! |  | 16 MeV | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
|  | 20 MeV | [7 empty] | 20 MeV | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! |  |  | 20 MeV | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! |  | 20 MeV | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
| 180 degrees | 6 MeV | [6 empty] | 180 degrees | 6 MeV | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! |  | 180 degrees | 6 MeV | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | 180 degrees | 6 MeV | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
|  | 9 MeV | [7 empty] | 9 MeV | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! |  |  | 9 MeV | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! |  | 9 MeV | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
|  | 12 MeV | [7 empty] | 12 MeV | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! |  |  | 12 MeV | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! |  | 12 MeV | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
|  | 16 MeV | [7 empty] | 16 MeV | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! |  |  | 16 MeV | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! |  | 16 MeV | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
|  | 20 MeV | [7 empty] | 20 MeV | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! |  |  | 20 MeV | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! |  | 20 MeV | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
| 270 degrees | 6 MeV | [6 empty] | 270 degrees | 6 MeV | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! |  | 270 degrees | 6 MeV | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | 270 degrees | 6 MeV | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
|  | 9 MeV | [7 empty] | 9 MeV | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! |  |  | 9 MeV | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! |  | 9 MeV | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
|  | 12 MeV | [7 empty] | 12 MeV | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! |  |  | 12 MeV | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! |  | 12 MeV | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
|  | 16 MeV | [7 empty] | 16 MeV | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! |  |  | 16 MeV | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! |  | 16 MeV | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
|  | 20 MeV | [7 empty] | 20 MeV | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! |  |  | 20 MeV | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! |  | 20 MeV | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!

## Mechanical
Mechanical |  |  |  |  | Functional
Error
2.1 Collimator rotation isocenter (TG142: 1mm from baseline)
2.2 Gantry rotation isocenter (TG142: 1mm from baseline)
2.3 Table rotation isocenter (TG142: 1mm from baseline)
In Machine QA/Service Mode, open Plans: C StarShot, G StarShot, and T StarShot
Orient Gafchromic film w/in Solid Water to capture StarShot exposures for respective plans.
100 SFD
Starshot analysis/images:
2.4 Coincidence of radiation and mechanical isocenter (TG142: 1mm from baseline)
In Machine QA/Service Mode, open daily Winston-Lutz RT Plan
Set WL phantom up w/lasers, image to fine tune position, and deliver.
WL Analysis:
2.5 Applicator interlocks (TG142: Functional)
| E- Cone (cm2) | Functionality | SRS Cone (mm2) | Functionality
| 6 | Functional | 4 | Functional
| 10 | Functional | 5 | Functional
| 15 | Functional | 7.5 | Functional
| 20 | Functional | 10 | Functional
| 25 | Functional | 12.5 | Functional
|  |  | 15 | Functional
|  |  | 17.5 | Functional
2.6 Table sag (TG142: 2mm from baseline)
Adjust table longitudinal to max displacement, towards gantry.
Tape Ruler to device on couch, such that coronal laser interseects with it.
Place at least 20cm of water equivalent at end of table (towards gantry).
Record displacement as measured by ruler.
| Displacement on Ruler (cm) | Baseline (cm)
| <1mm | <1mm
2.7 Table angle (TG142: 1o)
Using graph paper/IsoCal phantom/Varian Green Board, align with crosshairs
at 100cm SSD, rotate couch to align given phantom w/ crosshairs at other angles.
Check T=0o by manually realigning with crosshairs, also. Record DRO.
| Manual Rotation (o) | DRO (o)
0
315
270
45
90
2.8 Table travel maximum range movement in all directions (TG142: 2mm)
| DRO (o) |  | Baseline
| Min. | Max. | Min. | Max.
Longitudinal |  |  | 15.5 | 160.5
Lateral |  |  | -24.5 | 24.5
Vertical |  |  | -57 | 40.52
2.9 6DoF Couch Corrections (TG142: 1o)
Couch at Long. = 140, Lat. = 0, Vert. = 0
| DRO (o) | Meas.
Pitch | 3
0
357
Roll | 3
0
357

## EDW
EDW
3.1 Check wedge angle for 60o, full field (TG142: 2%)
3.2 Spot check for intermediate angle, field size (TG142: 2%)
3.3 Off-center ratios @ 80% field width @ 10cm (TG142: 2%)
IC Profiler, 9cm solid water + 0.9cm internal buildup = 9.9cm buildup. 100 SSD, 200 MU, 20x20 cm2
IC Profiler reported Y Axes Flatness/Symmetry @80% field width compared to measurements taken 6/2/2016
|  | EDW profiler files: |  | R:\Physics Documents\TrueBeam\Annual QA\2019\IC Profiler\Wedge Plate
| Energy | EDW X IN | 2016 Y Flatness (%) | Measured Y Flatness (%) | % Diff | 2016 Y Symmetry (%) | Measured Y Symmetry (%) | % Diff | Energy | EDW X IN | 2016 Y Flatness (%) | Measured Y Flatness (%) | % Diff | 2016 Y Symmetry (%) | Measured Y Symmetry (%) | % Diff | Energy | EDW X IN | 2016 Y Flatness (%) | Measured Y Flatness (%) | % Diff | 2016 Y Symmetry (%) | Measured Y Symmetry (%) | % Diff
| 6 MV | 10 | 7.5 |  |  | 7.1 |  |  | 10 MV | 10 | 6.5 |  |  | 6.3 |  |  | 18 MV | 10 | 5.8 |  |  | 5.5
|  | 15 | 10.8 |  |  | 10.8 |  |  |  | 15 | 9.4 |  |  | 9.5 |  |  |  | 15 | 8.5 |  |  | 8.4
|  | 20 | 14.5 |  |  | 14.7 |  |  |  | 20 | 12.6 |  |  | 12.9 |  |  |  | 20 | 11.4 |  |  | 11.5
|  | 25 | 18.3 |  |  | 18.8 |  |  |  | 25 | 16 |  |  | 16.5 |  |  |  | 25 | 14.4 |  |  | 14.7
|  | 30 | 22.3 |  |  | 23.1 |  |  |  | 30 | 19.5 |  |  | 20.3 |  |  |  | 30 | 17.6 |  |  | 18.1
|  | 45 | 30 |  |  | 32.6 |  |  |  | 45 | 29.4 |  |  | 31.6 |  |  |  | 45 | 28.2 |  |  | 30.1
|  | 60 | 29.6 |  |  | 31.6 |  |  |  | 60 | 29.8 |  |  | 31.7 |  |  |  | 60 | 30.1 |  |  | 32

## MLC
MLC
4.1 MLC transmission, all energies (TG142: 0.5% from baseline)
In Service/Machine QA Mode, open Plan: MLC Trans>Clear Couch Values, Manually retract leaves for Open fields
Farmer Chamber, 100 SSD, 5cm Depth
Orient Farmer Chamber perpendicular to MLC Leaves, such that interleaf/whole leaf transmission is averaged.
| 6 MV | Reading 1 | Reading 2 | Reading 3 | Average | Leakage (%) | 2018 Meas. DLG
| Open |  |  |  | #DIV/0! | #DIV/0! | 1.4%
| X1 Bank |  |  |  | #DIV/0!
| X2 Bank |  |  |  | #DIV/0!
| 6 FFF | Reading 1 | Reading 2 | Reading 3 | Average | Leakage (%) | 2018 Meas. DLG
| Open |  |  |  | #DIV/0! | #DIV/0! | 1.16%
| X1 Bank |  |  |  | #DIV/0!
| X2 Bank |  |  |  | #DIV/0!
| 10 MV | Reading 1 | Reading 2 | Reading 3 | Average | Leakage (%) | 2018 Meas. DLG
| Open |  |  |  | #DIV/0! | #DIV/0! | 1.68%
| X1 Bank |  |  |  | #DIV/0!
| X2 Bank |  |  |  | #DIV/0!
| 10 FFF | Reading 1 | Reading 2 | Reading 3 | Average | Leakage (%) | 2018 Meas. DLG
| Open |  |  |  | #DIV/0! | #DIV/0! | 1.47%
| X1 Bank |  |  |  | #DIV/0!
| X2 Bank |  |  |  | #DIV/0!
| 18 MV | Reading 1 | Reading 2 | Reading 3 | Average | Leakage (%) | 2018 Meas. DLG
| Open |  |  |  | #DIV/0! | #DIV/0! | 1.52%
| X1 Bank |  |  |  | #DIV/0!
| X2 Bank |  |  |  | #DIV/0!
4.2 Leaf position repeatability (TG142: 1mm)
In Service/Machine QA Mode, open RT Plan: MLC Repeatability
Position film underneath buildup, deliver once. Change position of MLC's, then re deliver field to same film, without moving the film
Film Scan:
4.3 MLC spoke shot (TG142: ≤1mm radius)
Spoke shot analysis/images:
4.4 Coincidence of light field and x-ray field, all energies (TG142: 2mm)
In Service/Machine QA Mode, open RT Plan: MLC Light Rad
Position film, and indicate light field. Deliver to film, beneath buildup.
Film Scan:
4.5 Dynamic deliviery (TG 142: See Below)
<0.35 cm max error RMS
95% of error counts <0.35 cm
Log File Analysis saved to DoseLab Database
Using DoseLab's FractionLab, analyze a few Trajectory Logs
Log File Analysis PDFs:

## Imaging
Imaging
Planar MV Imaging
5.1 Full range of travel SDD (TG142: 5mm)
Take images at the listed MVD positions and confirm measured field size agrees with set field size.
Technique: Highres
| MV Imaging Panel Position | Field Size (cm) |  | Measured Dimensions (cm)
| 0 | X= | 19 | X=
|  | Y= | 19 | Y=
| -20 | X= | 15 | X=
|  | Y= | 15 | Y=
| -40 | X= | 12 | X=
|  | Y= | 12 | Y=
| -60 | X= | 11 | X=
|  | Y= | 11 | Y=
5.2 Imaging dose (TG142: Baseline)
Using Buildup Cap, mount Farmer Chamber at Iso (similar to Output v. Gantry)
In Service mode, navigate to XVI tab, select MV>Acquisition>High Quality Single
10x10 cm2 | 100 SDD
Planar kV imaging
5.3 Beam quality/energy (TG142: Baseline)
5.4 Imaging dose (TG142: Baseline)
Using Buildup Cap, mount Farmer Chamber at Iso (similar to Output v. Gantry)
In Service mode, navigate to XVI tab, select: kV>Acquisition> ???
CBCT
5.5 Imaging dose (TG142: Baseline)
Using Buildup Cap, mount Farmer Chamber at Iso (similar to Output v. Gantry)
In Service mode, navigate to PVA tab, select ???

## Cross Cal
Chamber Cross Cal
| Setup conditions |  |  |  | Date:
| Material | Solid water
| Depth | 5 cm |  |  | Temp (C) | 22
| Backscatter | 11 cm |  |  | Pressure (mmHg) | 760
| Field size | 10x10 cm2 |  |  | PTP | 1
| SSD | 100
| MU | 100
| Polarity | +300V
|  |  | Chamber 1: |  | Chamber 2:
|  |  | N30013 SN9244 |  | S/N 1555 |  | Ratio
| Electrometer 1: | SuperMax | Reading 1 (nC) |  | Reading 1 (nC)
|  |  | Reading 2 (nC) |  | Reading 2 (nC)
|  |  | Reading 3 (nC) |  | Reading 3 (nC)
|  |  | Corrected Avg |  | Corrected Avg
| Electrometer 2: | Max4000 | Reading 1 (nC) |  | Reading 1 (nC)
|  |  | Reading 2 (nC) |  | Reading 2 (nC)
|  |  | Reading 3 (nC) |  | Reading 3 (nC)
|  |  | Corrected Avg |  | Corrected Avg
Barometer Cross Cal
[6 empty] | Date: | 2025-03-01
| Item | Value | Unit | Notes |  | Time: | 12:04
| Airport readout time | 11:52 |  | Omaha - Eppley
| Airport pressure | 30.3 | inMg | At sea level
| Airport pressure | 769.62 | mmMg | At sea level
| Airport pressure | 742.9022 | mmMg | At station
| Airport altitude | 297 | meters | Google Earth
| Clinic | Estabrook |  | <--choose
| Clinic altitude | 327 | meters | Google Earth
| Clinic pressure | 740.2458 | mmHg
| Device 1 pressure | 740.2 | mmHg | Your barometer / iPhone
| Device 2 pressure | 738.9 | mmHg | Mensor DPG-2400 by GK
| Difference 1 | -0.0061843 | % | % diff in P_TP or output
| Difference 2 | -0.1818 | % | % diff in P_TP or output
Google earth altittudes
| Airport | 297 | meters
| Estabrook | 327 | meters
| Jennie | 325 | meters
| Fremont | 361 | meters
Thermometer Cross Cal
[6 empty] | Date: | 2025-02-28
| Item | Value | Unit | Notes |  | Time: | 13:48
| Thermometer 1 | Traceable |  | SN240733782, Due 1/10/26
| Thermometer 2 | RadioShack |  | SN 06A01
| Temperature 1 | 20.4 | C
| Temperature 2 | 20.9 | C
| Temperature 1 | 293.55 | K
| Temperature 2 | 294.05 | K
| Difference | -0.17033 | % | % diff in P_TP or output

## TG-51 Photon
Annual QA Nebraska Methodist Hospital - Varian Truebeam 3348 | [5 empty] | Photon Correction Factors
Site Data | [6 empty] | 6 | 6 | 10 | 10 | 18
Physicist: | WSF | WSF | WSF | WSF | WSF | Reference: M(+300V) | 300 | 300 | 300 | 300 | 300
Date: | 2025-03-01 | 2025-03-01 | 2025-03-01 | 2025-03-01 | 2025-03-01 |  | 12.24 | 11.63 | 13.73 | 13.27 | 15.05
Linac Make/Model: | Varian Truebeam | Varian Truebeam | Varian Truebeam | Varian Truebeam | Varian Truebeam |  | 12.25 | 11.62 | 13.73 | 13.27 | 15.05
Linac Serial Number: | 3348 | 3348 | 3348 | 3348 | 3348
Nominal Photon Energy (MV): | 6 | 6 | 10 | 10 | 18 | Average | 12.245 | 11.625 | 13.73 | 13.27 | 15.05
Instrumentation Data | [5 empty] | Low: M(+150V) | 150 | 150 | 150 | 150 | 150
Chamber Model: | N30013 | N30013 | N30013 | N30013 | N30013 |  | 12.21 | 11.56 | 13.68 | 13.14 | 14.99
Chamber Serial Number: | 9244 | 9244 | 9244 | 9244 | 9244
Cavity Inner Radius (cm): | 0.305 | 0.305 | 0.305 | 0.305 | 0.305
ND,W(60Co) Gy/C: | 53920000 | 53920000 | 53920000 | 53920000 | 53920000 | Average | 12.21 | 11.56 | 13.68 | 13.14 | 14.99
Date of Calibration: | 2024-08-27 | 2024-08-27 | 2024-08-27 | 2024-08-27 | 2024-08-27 | Reverse: M(-300V) | 12.23 | 11.61 | 13.71 | 13.24 | 15.01
Electrometer Model: | Supermax | Supermax | Supermax | Supermax | Supermax
Electrometer Serial Number: | R172001 | R172001 | R172001 | R172001 | R172001
Pelec: | 0.999 | 0.999 | 0.999 | 0.999 | 0.999 | Average | 12.23 | 11.61 | 13.71 | 13.24 | 15.01
Date of Calibration: | 2023-10-02 | 2023-10-02 | 2023-10-02 | 2023-10-02 | 2023-10-02
Measurement Conditions
Setup (SSD or SAD): | SSD | SSD | SSD | SSD | SSD | Pressure Conversion: |  | 760 mmHg = 101.33 kPa
Field Size (cm2): | 10 | 10 | 10 | 10 | 10
Number of MU: | 100 | 100 | 100 | 100 | 100
Beam Quality | [5 empty] | kQ fit (TG51 update)
%dd(10): | 67.1107 | 64.0934 | 73.9308 | 71.0579 | 79.8643 | Chamber: | N30013
%dd(10)Pb (lead at 30cm +/- 1cm) | [5 empty] | A | 0.9652
%dd(10)x: | 67.1107 | 64.0934 | 73.9308 | 71.0579 | 81.188 | B | 2.141
kQ (NOTE: 63<PDD10x<86): | 0.99075 | 0.99467 | 0.98012 | 0.98489 | 0.96613 | C | -2.623
Output
Temperature (°C): | 19.6 | 19.6 | 19.6 | 19.6 | 19.6
Pressure (mmHg): | 740.1 | 740.1 | 740.1 | 740.1 | 740.1
Operating Voltage: | 300 | 300 | 300 | 300 | 300
PTP: | 1.0185 | 1.0185 | 1.0185 | 1.0185 | 1.0185
Ppol: | 0.99939 | 0.99935 | 0.99927 | 0.99887 | 0.99867
Pion: | 1.0029 | 1.0057 | 1.0037 | 1.01 | 1.004
Mraw (nC): | 12.24 | 11.63 | 13.73 | 13.27 | 15.05
| 12.25 | 11.62 | 13.73 | 13.27 | 15.05
| 12.25 | 11.63 | 13.72 | 13.27 | 15.06
Corrected Ion Chamber Rdg, M (nC): | 12.4894 | 11.8896 | 14.0082 | 13.622 | 15.3582
Dose (cGy) to Water at 10 cm depth: | 66.7199 | 63.7671 | 74.0305 | 72.3403 | 80.0065
Clinical %dd(10): | 66.4 | 63.5 | 73.5 | 71.1 | 79.5
Dose (cGy) / MU at dmax: | 1.0048 | 1.0042 | 1.0072 | 1.0174 | 1.0064
Post Mraw (nC): | 12.19 | 11.58 | 13.63 | 13.04 | 14.965
Post Corrected Ion Rdg, M (nC): | 12.4316 | 11.8419 | 13.9095 | 13.3859 | 15.2681
Post Dose to Water at dref (cGy): | 66.4111 | 63.5112 | 73.5092 | 71.0865 | 79.537
Post Dose / MU at dmax: | 1.0002 | 1.0002 | 1.0001 | 0.99981 | 1.0005
M = Pion*PTP*Pelec*Ppol*Mraw
Dose to Water at 10 cm depth = M*kQ*NDW(60Co)

## TG-51 Electron
Annual QA Nebraska Methodist Hospital - Varian Truebeam 3348 | [5 empty] | Electron Correction Factors
Site Data | [6 empty] | 6 | 9 | 12 | 16 | 20
Physicist: | WSF | WSF | WSF | WSF | WSF | dref | 1.3106 | 2.0652 | 2.9461 | 3.9679 | 4.9482
Date: | 2025-03-01 | 2025-03-01 | 2025-03-01 | 2025-03-01 | 2025-03-01 | Reference: M(+300V) | 300 | 300 | 300 | 300 | 300
Linac Make/Model: | Varian Truebeam | Varian Truebeam | Varian Truebeam | Varian Truebeam | Varian Truebeam |  | 19.7 | 19.87 | 20.09 | 20.05 | 19.55
Linac Serial Number: | 3348 | 3348 | 3348 | 3348 | 3348 |  | 19.71 | 19.89 | 20.09 | 20.03 | 19.55
Nominal Electron Energy (MeV): | 6 | 9 | 12 | 16 | 20
Instrumentation Data | [5 empty] | Average | 19.705 | 19.88 | 20.09 | 20.04 | 19.55
Chamber Model: | N30013 | N30013 | N30013 | N30013 | N30013 | Low: M(+150V) | 150 | 150 | 150 | 150 | 150
Chamber Serial Number: | 9244 | 9244 | 9244 | 9244 | 9244 |  | 19.48 | 19.65 | 19.85 | 19.8 | 19.32
Cavity Inner Radius (cm): | 0.305 | 0.305 | 0.305 | 0.305 | 0.305 |  | 19.48 | 19.66
ND,W(60Co) Gy/C: | 53990000 | 53990000 | 53990000 | 53990000 | 53990000
Date of Calibration: | 2024-08-27 | 2024-08-27 | 2024-08-27 | 2024-08-27 | 2024-08-27 | Average | 19.48 | 19.655 | 19.85 | 19.8 | 19.32
Electrometer Model: | Supermax | Supermax | Supermax | Supermax | Supermax | Reverse: M(-300V) | 19.65 | 19.83 | 20.03 | 20 | 19.52
Electrometer Serial Number: | R172001 | R172001 | R172001 | R172001 | R172001 |  | 19.63 | 19.82
Pelec: | 0.999 | 0.999 | 0.999 | 0.999 | 0.999
Date of Calibration: | 2023-10-02 | 2023-10-02 | 2023-10-02 | 2023-10-02 | 2023-10-02 | Average | 19.64 | 19.825 | 20.03 | 20 | 19.52
Measurement Conditions | [5 empty] | dref+0.5rcav | 1.4631 | 2.2177 | 3.0986 | 4.1204 | 5.1007
Distance SSD: | SSD | SSD | SSD | SSD | SSD | M(dref+0.5rcav) | 19.65 | 19.85 | 20.01 | 19.93 | 19.43
Field Size on Surface (cm2): | 10 | 10 | 10 | 10 | 10 |  | 19.64 | 19.84
Number of MU: | 100 | 100 | 100 | 100 | 100
Beam Quality | [5 empty] | Average | 19.645 | 19.845 | 20.01 | 19.93 | 19.43
I50: | 2.343 | 3.5652 | 4.9921 | 6.6471 | 8.2348
R50: | 2.351 | 3.6086 | 5.0769 | 6.7799 | 8.4136
Reference Depth, dref : | 1.3106 | 2.0652 | 2.9461 | 3.9679 | 4.9482
kecal: | 0.897 | 0.897 | 0.897 | 0.897 | 0.897 | Pressure Conversion: |  | 760 mmHg = 101.33 kPa
k'R50: | 1.0279 | 1.0171 | 1.0083 | 1.0017 | 0.99767
Output
Temperature (°C): | 19.65 | 19.65 | 19.65 | 19.65 | 19.65
Pressure (mmHg): | 740.2 | 740.2 | 740.2 | 740.2 | 740.2
Operating Voltage: | 300 | 300 | 300 | 300 | 300
PTP: | 1.0186 | 1.0186 | 1.0186 | 1.0186 | 1.0186
Ppol: | 0.99835 | 0.99862 | 0.99851 | 0.999 | 0.99923
Pion: | 1.0117 | 1.0116 | 1.0122 | 1.0123 | 1.012
PgrQ(cyl): | 0.99696 | 0.99824 | 0.99602 | 0.99451 | 0.99386
Mraw (nC): | 19.71 | 19.87 | 20.09 | 20.05 | 19.55
| 19.71 | 19.89 | 20.09 | 20.03 | 19.55
|  | 19.87 | 20.08 | 20.02 | 19.56
Corrected Ion Rdg, M (nC): | 20.2569 | 20.4316 | 20.6586 | 20.6146 | 20.1209
Dose to Water at dref (cGy): | 100.534 | 100.4589 | 100.4765 | 99.4546 | 96.62
Clinical %dd(dref): | 100 | 100 | 99.8 | 98.7 | 95.6
Dose / MU at dmax: | 1.0053 | 1.0046 | 1.0068 | 1.0076 | 1.0107
Post Mraw (nC): | 19.615 | 19.79 | 19.95 | 19.89 | 19.345
Post Corrected Ion Rdg, M (nC): | 20.1593 | 20.3425 | 20.518 | 20.4671 | 19.9065
Post Dose to Water at dref (cGy): | 100.0494 | 100.0209 | 99.7928 | 98.743 | 95.5905
Post Dose / MU at dmax: | 1.0005 | 1.0002 | 0.99993 | 1.0004 | 0.9999
M = Pion*PTP*Pelec*Ppol*Mraw
Dose to Water at 10 cm depth = M*Pgr*k'R50*kecal*NDW(60Co)

## PDDs
PDDs
| Date | 2025-03-01
| Phantom | 30x30x30 water
| Chamber | A16 micro
| Vol | 0.007
| NDw | 5040000000 | Gy/C |  | Photon PDD10cm (%) (1%) |  |  |  | Electron R50 (cm) (1mm)
| Cal | 2023-09-28 |  |  | Energy | Comm | TPS | 2025 | Energy | Comm | TPS | 2025 | 2025 I50
| Rcav | 0.12 | cm |  | 6 MV | 66.5 | 66.29 | 67.1107 | 6 MeV | 2.34 | 2.35 | 2.351 | 2.343
| 0.6Rcav | 0.072 | cm |  | 6 FFF | 63.2 | 63.45 | 64.0934 | 9 MeV | 3.58 | 3.57 | 3.6086 | 3.5652
| 0.5Rcav | 0.06 | cm |  | 10 MV | 73.3 | 73.37 | 73.9308 | 12 MeV | 5.03 | 4.98 | 5.0769 | 4.9921
| Range | Low |  |  | 10 FFF | 71 | 70.97 | 71.0579 | 16 MeV | 6.7 | 6.61 | 6.7799 | 6.6471
| Units | pC |  |  | 18 MV | 79.6 | 79.58 | 79.8643 | 20 MeV | 8.33 | 8.25 | 8.4136 | 8.2348
| Leakage | 0.001 pA
| 6 MV |  |  |  | 10 MV |  |  |  | 6 FFF |  |  |  | 10 FFF |  |  |  | 18 MV
| POM | EPOM | Value | PDD | POM | EPOM | Value | PDD | POM | EPOM | Value | PDD | POM | EPOM | Value | PDD | POM | EPOM | Value | PDD
| 1.5 | 1.428 | 195.1 | 99.9488 | 2.3 | 2.228 | 196.3 | 99.9491 | 1.3 | 1.228 | 196.5 | 99.7969 | 2.1 | 2.028 | 200.1 | 99.8503 | 3 | 2.928 | 198.1 | 99.5978
| 1.7 | 1.628 | 194.8 | 99.7951 | 2.4 | 2.328 | 196.4 | 100 | 1.2 | 1.128 | 196.2 | 99.6445 | 2.2 | 2.128 | 200.2 | 99.9002 | 3.1 | 3.028 | 198.5 | 99.7989
| 1.6 | 1.528 | 195.1 | 99.9488 | 2.5 | 2.428 | 196.3 | 99.9491 | 1.4 | 1.328 | 196.9 | 100 | 2.3 | 2.228 | 200.4 | 100 | 3.2 | 3.128 | 198.4 | 99.7486
| 1.4 | 1.328 | 195.2 | 100 | 2.6 | 2.528 | 196.2 | 99.8982 | 1.5 | 1.428 | 196.4 | 99.7461 | 2.4 | 2.328 | 200.4 | 100 | 3.3 | 3.228 | 198.9 | 100
| 1.3 | 1.228 | 194.8 | 99.7951 | 2.7 | 2.628 | 195.8 | 99.6945 | 1.6 | 1.528 | 196.5 | 99.7969 | 2.5 | 2.428 | 200 | 99.8004 | 3.4 | 3.328 | 198.7 | 99.8994
| 5.07 | 4.998 | 168.4 | 86.2705 | 2.8 | 2.728 | 195.4 | 99.4908 | 1.7 | 1.628 | 196.1 | 99.5937 | 2.6 | 2.528 | 199.5 | 99.5509 | 3.5 | 3.428 | 198.6 | 99.8492
| 10.07 | 9.998 | 131 | 67.1107 | 2.2 | 2.128 | 196.2 | 99.8982 | 5.07 | 4.998 | 167.3 | 84.967 | 5.07 | 4.998 | 181.2 | 90.4192 | 5.07 | 4.998 | 191.75 | 96.4052
[5 empty] | 5.07 | 4.998 | 180 | 91.6497 | 10.07 | 9.998 | 126.2 | 64.0934 | 10.07 | 9.998 | 142.4 | 71.0579 | 10.07 | 9.998 | 158.85 | 79.8643
[5 empty] | 10.07 | 9.998 | 145.2 | 73.9308 |  |  |  | 0 | [7 empty] | 0
| 6 MeV |  |  |  | 9 MeV |  |  |  | 12 MeV |  |  |  | 16 MeV |  |  |  | 20 MeV
| POM | EPOM | Value | Ratio | POM | EPOM | Value | Ratio | POM | EPOM | Value | Ratio | POM | EPOM | Value | Ratio | POM | EPOM | Value | Ratio
| 1.1 | 1.04 | 206.9 | 0.97 | 2 | 1.94 | 210.9 | 0.99811 | 2.5 | 2.44 | 210.9 | 0.99387 | 2.8 | 2.74 | 215.3 | 0.99815 | 1.8 | 1.74 | 220.7 | 0.99819
| 1.2 | 1.14 | 210.6 | 0.98734 | 2.1 | 2.04 | 211.3 | 1 | 2.6 | 2.54 | 211.7 | 0.99764 | 3 | 2.94 | 215 | 0.99675 | 2.2 | 2.14 | 220.5 | 0.99729
| 1.3 | 1.24 | 212.8 | 0.99766 | 2.2 | 2.14 | 210.9 | 0.99811 | 2.7 | 2.64 | 212 | 0.99906 | 3.2 | 3.14 | 214.9 | 0.99629 | 2.6 | 2.54 | 219.7 | 0.99367
| 1.4 | 1.34 | 213.3 | 1 | 2.3 | 2.24 | 210.5 | 0.99621 | 2.8 | 2.74 | 212.1 | 0.99953 | 2.6 | 2.54 | 215.6 | 0.99954 | 1.4 | 1.34 | 221.1 | 1
| 1.5 | 1.44 | 212.1 | 0.99437 |  |  |  |  | 2.9 | 2.84 | 212.2 | 1 | 2.4 | 2.34 | 215.7 | 1 | 1.2 | 1.14 | 221.1 | 1
| 2.3 | 2.24 | 125.7 | 0.58931 | 3.5 | 3.44 | 122.2 | 0.57832 | 3 | 2.94 | 211.9 | 0.99859 | 2.2 | 2.14 | 215.7 | 1 | 2 | 1.94 | 221.1 | 1
| 2.4 | 2.34 | 107.2 | 0.50258 | 3.6 | 3.54 | 109 | 0.51585 | 4.9 | 4.84 | 121.4 | 0.5721 | 6.6 | 6.54 | 115.4 | 0.535 |  |  |  | 0
| 2.5 | 2.44 | 89.08 | 0.41763 | 3.7 | 3.64 | 95.72 | 0.45301 | 5 | 4.94 | 111 | 0.52309 | 6.7 | 6.64 | 108.4 | 0.50255 | 8.2 | 8.14 | 115.1 | 0.52058
[9 empty] | 5.1 | 5.04 | 101.6 | 0.47879 | 6.8 | 6.74 | 100.7 | 0.46685 | 8.3 | 8.24 | 110.3 | 0.49887